## Test Report
Test Inlabtec Serial Diluter
Company/ Lab | iNLABTEC AG/ R&D
Date | 2026-09-15
Inlabtec Serial Diluter Serial No. | 2024005
Examiner | XP
Test reason | System check
Serological Pipette
Type/ Manufacturer | 94010/ TPP
LOT Nr. | 20110091
Inlabtec Serial Dilution Bags
LOT Nr. | PE520447
1 ml Pipette Tips (not relevant for Serial Diluter UA/UC)
Type/ Manufacturer
LOT Nr.
Test Liquid
z-factor (reciprocal density) [ml/g] | 1.003
Water, z-factor = 1.003 Peptone saline diluent, z-factor = 0.993
Weight Serial Dilution Bags
Total weight of 6 Serial Dilution Bags [g] | 1.506
Average weight per bag [g] | 0.251
Measurement Serial Dilution Bag | Total Weight [g] | Weight Liquid [g] | Liquid Volume [ml]
1 | 9.345 | 9.094 | 9.1213
2 | 9.286 | 9.035 | 9.0621
3 | 9.25 | 8.999 | 9.026
4 | 9.329 | 9.078 | 9.1052
5 | 9.263 | 9.012 | 9.039
6 | 9.261 | 9.01 | 9.037
| Arithmetic average [ml] |  | 9.0651
| Set value [ml] |  | 9
| Deviation from the set value [ml] |  | 0.065114
| Relative deviation from the set value R (Inaccuracy) |  | 0.7235%
| Standard deviation [ml] |  | 0.039427
| Relative standard deviation P (Precision) |  | 0.4349%
| Max. error according to ISO 7218:2024, 6.4.4.4 |  | 2%
| Determined maximum error |  | 1.593%
| Test result |  | Passed